# ANNEX II + III :	 TECHNICAL SPECIFICATIONS + TECHNICAL OFFER

- Contract title: Supply of ambulance vehicles and medical equipment – Lot 2	p 1 /…

Publication reference: RORS 284/Cityof Kikinda/TD2-relaunch

Columns 1-2 should be completed by the Contracting Authority

Columns 3-4 should be completed by the tenderer

Column 5 is reserved for the evaluation committee

Annex III - the Contractor's technical offer

The tenderers are requested to complete the template on the next pages:

- Column 2 is completed by the Contracting Authority shows the required specifications (not to be modified by the tenderer),

- Column 3 is to be filled in by the tenderer and must detail what is offered (for example the words “compliant” or “yes” are not sufficient)

- Column 4 allows the tenderer to make comments on its proposed supply and to make eventual references to the documentation

The eventual documentation supplied should clearly indicate (highlight, mark) the models offered and the options included, if any, so that the evaluators can see the exact configuration. Offers that do not permit to identify precisely the models and the specifications may be rejected by the evaluation committee.

The offer must be clear enough to allow the evaluators to make an easy comparison between the requested specifications and the offered specifications.

| 1. Item Number | 2. Specifications Required | 3. Specifications Offered | 4. Notes, remarks, ref to documentation | 5. Evaluation Committee’s notes |
| --- | --- | --- | --- | --- |
| 1. | Defibrillator (1 pcs) General requirements: - manual biphasic defibrillator/monitor with pedals and pads - defibrillation modes (min): manual, AED - synchronous cardioversion included - the device meets EN 1789 standard or equivalent - the device meets the standard IP 44 or higher - weight: maximum 8.5 kg with battery/batteries and printer paper - Power supply: AC, DC and battery supply - device operable on AC 220V -device operable from vehicle power supply connection 12V DC - battery type: rechargeable lithium – ion with capacity of min. 4.5 hours of ECG monitoring - full functionality of the device with or without a battery when the device is connected to an external AC/DC power supply - operability of the device at a temperature of minimum 0 degrees Celsius - operability of the device at a temperature of min. -5 degrees of Celsius lasting min. 20 minutes Data monitoring (min.): ECG, NIBP, pulse oximetry (SpO2), capnography (CO2), temperature, IBP ECG - the possibility of recording a 12-channel ECG with automatic analysis NIBP - the possibility of measuring NIBP using the oscillometric method pulse oximetry (SpO2) - the possibility of measuring SpO2 with the possibility of upgrading the SpCO and SpMetH module capnography (CO2) - the possibility of measuring CO2 in intubated and non-intubated patients TEMPERATURE - the possibility of temperature measurement through disposable skin sensors DISPLAY OF DATA: - monitor: colour screen, 6.5” diagonal minimum, high contrast and operable under strong sunlight; display of at least 3 curves and numerical data - thermal printer with transparent parallel printing of at least 3 channels PADLES AND PADS: - adults and pediatric paddles with the possibility of defibrillation using disposable self-adhesive electrodes - controls for charging and delivering electric shocks on the defibrillator's pedal CHARGE TIME: - within 10 seconds electric charge time to maximum energy level PACING: - external cardiac pacemaker with the possibility of sending "fixed" or "on demand" stimulation pulses OTHER: - device possess sound metronome with appropriate protocol of CPR - the possibility of data transmission via USB connection to a computer - accessories: ECG patient cable for monitoring and diagnostics (1 pcs), SpO2 finger sensor, reusable (1 pcs), connection cable for SpO2 sensor (1 pcs), blood pressure cuff for adults with cuff connection hose (1 pcs), pair of self-adhesive adult pacing/defibrillation electrodes (1 pcs), defibrillation gel (1 pcs), printer paper (1 pcs), docking station for placement and charging of the device in the vehicle, power cord (1 pcs), standard device bag for accessories designed for emergency services (1 pcs), - device lithium-ion battery/batteries (quantity depending of the device) delivered |  |  |  |
| 2. | Mobile battery charger (1 pcs) – compatible to defibrillator device from point 1 |  |  |  |
| 3. | 12-channel ECG device (1pcs) - 3-channel, AC/DC-battery EKG device - integrated power/battery supply - 12-channel colour display screen (diagonal min. 8 inches) with the possibility of display of all 12 channels - the device has integrated filters for stabilizing the ECG signal - the device has a system for detecting misplaced electrodes - the device has a quality check of each lead individually, displayed in different colors depending on the status, the so-called “color-coded” check - the device has a program for interpretation of adult and pediatric ECG - the device must support the formula for QT correction (min): Bazett, Frederica, Fremingham, Hadges - memory of min. 100 cases (ECG recordings) - the possibility of exporting the recording to a USB memory (min. PDF export format) - connectivity: the device must possess LAN and WiFi module - internal 12-channel thermal printer for printing on A4 thermal paper with the possibility of reviewing the ECG before printing, exporting and/or saving; speed options min. 12,5-25-50 mm/s - potted, rubber or silicone alpha numeric keypad, with direct function buttons - full charged battery duration: min. 7 hours - net weight: max. 5kg - delivered with standard accessories |  |  |  |
| 4. | Handheld, portable ultrasound machine with display device, licenced software, battery charger, 3 probes (1 pcs) - operating modes (min): B, M, Colour flow map, Pulse Wave Doppler - the possibility of pre-set parameters and measurements for areas of application (min): abdominal, cardiac adult, pediatric cardiac, lung, TCD - the probe has “+” and “-“ buttons (adjustment of doppler gain, depth, focus), as well as a button to freeze the image - display device compatible with the probe (diagonal min. 9.6 inches) - licenced software - the possibility of zooming the image with two fingers by dragging on the screen of the display device - the possibility of installing an application for remote ultrasound diagnostics in real time - the possibility of connecting the probe to several devices (tablet, mobile phone) - DICOM connection and wireless data transfer - handheld bag for the device |  |  |  |
| 5. | Chest compression system (1 pcs) - automated, portable, battery-powered (and AC/DC) mechanical chest compressor - adjustable therapy modes/ventilation cycles (min): 30:2 /continuous - frequency of compressions capability (min.): 100/min - device is operabile during defibrilation (compatible with defibrilator device) - adjustability in accordance with the patient's chest height - no restrictions on the weight of the patient - device has a system for patient position stabilization during operation (belts, straps or equvivalent) - operability at the envirnoment temperature (min.): 0°C to +40°C; - min. 20 minutes operability at the temperature from – 20°C to +50°C - casing/ protection class (min): IP 33 - weight of the device (max): 11kg - battery charge indicator on control panel - operability of the battery (min.): 40 minutes of continuous operation - ventilation cycles choice on control panel - battery charger included - min. 1 (one) batery included - power cables: both AC and DC included - transport bag or suitcase included |  |  |  |
| 6. | Skill Trainer simulator (1 set): manikin, data display device, control device, software - skill trainer simulator for CPR, artificial ventilation, intubation, IV Injections - manikin: adult, full body Airway features: - Anatomically accurate, realistic airway - Head tilt / chin lift - Jaw thrust - BVM ventilation - Oral or nasal intubation - Realistic chest rise and fall - Sellick maneuver CPR and cardiac features: - Chest compression - Compression measurement and feedback - Defibrillation (manual and automatic) - ECG monitoring - Pulse monitoring Vascular Features - Multivenous IV arm simulation of IV drug administration, IV insertion, infusion - Adjustable blood pressure levels (Systolic/Diastolic) Simulation and educational features: - Creation and editing of events scenarios - Pre-programmed scenarios - CPR performance evaluation System accessories: - patient/manikin simulator monitor for vital signs - wireless control panel / device - USB port and cable for data transfer - adapter for hard paddles - transport case - artificial blood - lubricant - tools - manual - operated on batteries or 220V power - appropriate software, with unrestricted time usage |  |  |  |
| 7. | Compact patient monitor (1 pcs) - monitor for continuously monitoring of physiological parameters (min.): ECG, resp, temp, SPO2 - display: min 12.1 inches diagonal - waveform display - delivered with standard accessories |  |  |  |
| 8. | Inhalator with nebuliser with PIF-Control (1 pcs) - inhalation time less than 5 min. - Peak Inspiratory Flow control - weight: max. 2kg - pressure: min. 1.6bar |  |  |  |
| 9. | Capnograph kit (1 pcs) - direct connection to endotracheal tubes, laryngeal masks or facial masks via the airway adapters (adult and pediatric) |  |  |  |
| 10. | Oxygen bottle with reduction valve (1 pcs) Volume: 10 litres |  |  |  |
| 11. | Non-Invasive Ventilation BiPAP Machine (1 pcs) - Ventilation Modes (min.): S, ST, PSV, CPAP – intended for adult and peadiatric >10kg - Displayed data: Ppeak, EPAP, Leakage, MV, Vte % in TgV, Total Rate I:E, Spont Rate, % spont, Inspiratory time, Rise time, SpO2, Pulse, Rate |  |  |  |
| 12. | Work uniform set (jacket, sweatshirt, pants, T-shirt) – 40 sets Jacket: - high visibility reflective bands night/day - material: water-repellent or waterproof - colour: red - winter lining included and removable - the sleeves removable - the hood removable - zipper included - two pockets minimum - the name of the institution printed; printed visibility requirement of Interreg-IPA Cross-border Cooperation Romania-Serbia Programme; Pants: - -high visibility reflective bands night/day - material: water-repellent or waterproof - trouser leg pockets - colour: red T – shirt - material: cotton - colour: red - the name of the institution printed Sweatshirt: - colour: red - the name of the institution printed - zipper included |  |  |  |
| 13. | Medicine cabinet (1 pcs) - two segments in one unit - two-winged cabinet - dimensions: min. 80cm/40cm/180cm - upper doors filled with glass - door locks installed |  |  |  |
| 14. | Emergency service bag (4 pcs) - padded handle - detachable shoulder straps - min. 3 (three) external pockets, with elastic bands - inner layout adaptable to the user’s needs via dividers - inner detachable compartments for classification of the material min. 3 (three) - high visibility reflective bands night/day & reflective pullers - reinforced bottom - waterproof - colour: red - load capacity: min. 10kg |  |  |  |
| 15. | Ampoule holder (4 pcs) - lightweight handhold ampule bag - foldable with zipper - weight (max.): 700gr - interior with min. two parts, with elastic bands to hold medium and small ampoules - capacity: min. 35 ampules |  |  |  |
| 16. | Therapy wheeled cart (1 pcs) - 3 floors/shelfs - four wheels installed - shelf stops - material resistant to disinfectants |  |  |  |
| 17. | Automatic intraosseous device – adult (2 pcs) - no exposed needle - no assembly required - no external power source or battery required -needle penetration depth: max. 45mm - safety mechanism installed |  |  |  |
| 18. | Automatic intraosseous device – pediatric (2 pcs) - no exposed needle - no assembly required - no external power source or battery required -adjustable needle penetration depth according to the age of the patient - safety mechanism installed |  |  |  |
| 19. | One-piece paravane (2 pcs) - white, opaque, cotton canvas - construction of metal pipes - dimensions: min. height 175cm, min. width 70cm - wheels installed |  |  |  |
| 20. | Pelvic Sling (1 pcs) - for stabilization of pelvic fractures with the correct force - reusable - resistant to washing and disinfection means - compatible with X-rays, CT scans, MRI (usage without folding during mentioned diagnostics) - tapered for comfort over abdomen |  |  |  |
| 21. | Flat Fold Splint (10 pcs) - immobilization of extremities - built from a thin core of aluminium alloy, sandwiched between two layers of closed-cell foam - can be bent into any of three simple curves - radiolucent - reusable |  |  |  |
| 22. | Critical Burn Kit (1 set) - set against burns - burn dressing for: face (min. 1 pcs), hand (min. 2 pcs), foot (min. 2 pcs) - one scissors - all burn dressing made of high percentage of water material - kit content folded in protective impervious caring bag/bags |  |  |  |
|  | Note: medical equipment must be certified/registered in accordance to the national “Law on Medical Devices” and “Rulebook on registration of medical device” of the Republic of Serbia; |  |  |  |